## Slide 1: Neurolabware collected wisdom
Whatever makes it work…

## Slide 2: General
Last group of digits in file name has to have three digits 000
sbxread_all() reads all frames in file and returns raw data and info struct
Orientation of Knobby/Lab frame axes in Scanbox images: see photo below. Note that this is HORRIBLE, because it means that there is now direct rotation that brings the image orientation in sync with any top-down view in the lab frame.
If the mouse is looking towards the computer, then the Scanbox images must be flipped vertically and then rotated 90deg right to orient mouse A direction towards up in image, and mouse L direction towards L in image.

## Slide 3: Mesoscopic panorama
  Set Magnification >= 2x. The larger the magnification the more even the brightness across columns in the panorama
  Choose no of patches such that calculated FOV indicator is not red
  E.g. for good coverage of most of the 5 x 4.5 mm FOV:

## Slide 4: Mesoscopic panorama + z stack
First set up panorama as on previous slide (disable rolling average)
Move to deepest position (here: -550 um)
Set up knobby scheduler as shown here, using range, step, frames inputs
Enable arm and return
Increase number of Total frames by one. Due to some bug, Scanbox always saves one frame less than requested
Grab
Adjust laser power manually during acquisition to avoid overexposure
Use function mesoscope_stack_to_tif() to convert

## Slide 5: Defining ROIs in mesoscope window
Click “Add ROI”
Position ROI
Double-clicking ROI adds it to Mesoscope panel

## Slide 6: Problem: stripe brightness mismatches in mesoscope panorama mode
Yang advises:
Observe beam at entrance port iris with IR viewer and center if necessary
Alternatively, walk beam during acquisition until brightness flatness improves

## Slide 7: Open Questions
What is that unlabelled slider for that appears in the meso window after adding an ROI?

## Slide 8: BUGS

## Slide 9
Scanbox always saves one frame less than specified in the “Total frames” textbox in the GUI. Not sure if first or last frame is omitted.

## Slide 10: Error after manual z move and pressing focus: (mesoscope panorama enabled)

## Slide 11: Bug: setting up a meso panorama before first acquisition

## Slide 12: Inconsistency:
Zero XYZ buttons in Scanbox seem to move objective (a warning is displayed), whereas Zero button on Knobby just “renames” the current position to all-zeros  make nomenclature consistent
A “rename to zero” button in the GUI would be useful

## Slide 13: BUG
Laser power sliders crashed at one point without an error message: change of slider pos would not cause change of laser power, displayed percentage not changing either

## Slide 14: BUG: meso window ignores flipping h axis: sbconfig.hsync_sign = 1; (% 0-normal, 1-flip horizontal) (orientation within strip is correct, but sequence of stripes would have to be flipped horizontally